## Slide 1: Survey-mode measurements and analysis
M. Floyd K. Palamartchouk
Massachusetts Institute of Technology Newcastle University
GAMIT-GLOBK course University of Bristol, UK 12–16 January 2015
Material from R. King, T. Herring, M. Floyd (MIT) and S. McClusky (now ANU)

## Slide 2: Measurement Strategies I: Occupation Time
Given time and personnel constraints, what are the trade-offs between between spatial and temporal density?
Ideally, you would like for the white-noise position uncertainty for an occupation to contribute to the velocity uncertainty at a level less than the usually dominant long-period correlated noise.
Typical white-noise uncertainties (Horizotnal and Vertical) as a function of occupation time:
6-8 hrs: 2-2.5 mm H, 5-10 mm V
12-24 hrs: 1.0-1.5 mm H, 3-5 mm V
36-48-hrs: 0.7-1.0 mm H, 2-4 mm V
Observations over 3 or more days will give you more redundancy
Observations of 5 or more days will be necessary for mm-level vertical uncertainties
If your region has few continuous stations, you should consider running one or two survey-mode stations for the entire time of the survey to provide continuity
Notes: If your scientific\ goals require a spatially dense network , then you may chose to rely on having two or more sites whose motions are expected to be the same and use this spatial redundancy as s substitute for temporal redundancy in identifying outliers. The spatial redundancy would be more important than the temporal if the error in a position determination is correlated over the occupation time (e.g. a single flawed setup continued for several days).

## Slide 3: Precision v session length for network processing
Notes: This slide shows repeatability in position over 31 days as a function session length of a test site in central Italy when processed with 2 to 16 continuous reference sites spanning 180-600 km. See Firuzabadi and King, GPS Solutions, doi:10/1007/s10291-0110218-8, 2011, for details of the study.

## Slide 4: Measurement Strategies II: Monuments and Instrumentation
Issues in site and antenna selection :
Monument stability
Accessibility
Ease of setup
Multipath
Vandalism
  • There is no clear prescription for all cases; let’s look at some examples

## Slide 5: Three primary mounting options
Tripod with optical or physical plummet
Spike mount
Mast
Courtesy UNAVO web page
Notes: Spike mount: Light-weight, fixed height, low profile and easy to set, but vulnerable to low-ground reflections and long-period multipath.
Tripod: Avoids low-ground obstructions and long-period multipath, but heavy, high-profile, and vulnerable to setup errors.
Tech 2000 mast: medium weight, fixed height, avoids low-ground obstructions and long-period multipath, but 1st-time setup time-consuming.

## Slide 6: Low mount in a good enviroment
STVP
Steven’s Pass, Cascades Range in western Washington
18-cm spike mount
No long-term repeatability yet, but 44 hrs
of observations in 2012 give formal uncertainties 0.5 mm horizontal, 3 mm
vertical. Note minimal long-period signal
Scattering.
Notes: This is an example of a good sGPS site observed with a low-mounted antenna. The ground drops offf in a 270-degree arc, trees block the view and generate reflections in a 90-degree arc, mostly in the north-polar “hole” of SV visibility.

## Slide 7: Low mount in a dirty environment
B059
Roadside meadow in western Washington
12.5-cm spike mount
Two 24-hr measurements in 2012 agree at 1 mm horizontal, 4 mm though formal uncertainties 2 mm, 10mm due to high random noise (diffuse multipath or water vapor?) Note minimal long-period signal
scattering. Long-term scatter 3 mm horizontal, 5 mm vertical (monument instablity?)

## Slide 8: High mount in a dirty environment
C033
Old survey mark in dirt in central Washington.
Tripod mount. (Train blockage was short-lived)
2012 19-hr session and 5-hr session agree at 1.5 mm horizontal, 3 mm vertiical. Long-term repeatability 2 mm horizontal, 12 mm vertical.
Surprisingly little short-period multipath (dry dirt?)

## Slide 9: Low mount on a slope
LYFR
Rocky river bank in eastern Oregon
12.5-cm spike mount
Single 14-hr session . Long-period multipath due to slope and/or reflective rocks ?
Notes: We have yet to determined the relationship between the height of the spike mount and the effect on long-period multipath. Is there a critically bad resonance height (remember that L1 has a wavelength of 19 cm, and L2 24 cm)? We’ve seen some evidence of this with pillars. Is there an optimal height that minimizes the long-period multipath? Possibly, but the curve may be shallow.

## Slide 10: Special Characteristics of Survey-mode Data
Editing is critical: every point counts
Usually combined with cGPS data to provide continuity and a tie to the ITRF
Appropriate relative weighting needed in combining with cGPS data
Antenna meta-data may be more complicated
Heights may be problematic if different antennas used
Seasonal errors behave differently than in cGPS data: best strategy is to observe at the same time of the year (unlike cGPS, which has minimal seasonal sensitivity at 1.5, 2.5, 3.5 ….years total span)

## Slide 11: Analysis Strategy
Generate time series and aggregated h-files for each survey
  Use spans less than ~ 30 days to avoid biasing the position estimate from an in correct velocity
  Include cGPS data only on days when sGPS data are available to maintain common-mode cancellation
  Strenghens the sGPS positions estimates within each survey and allows better assessment of the long-term statistics
  Edit carefully the daily values within each span
Generate time series and a velocity using the aggregated h-files from a span of 3 or more years
  Edit carefully the long-term time series
  Add 0.5 of white noise (sig_neu) to the cGPS estimates from each span to avoid overweighting the cGPS position estimates
  Use a separate (ie.g PBO) analysis of the daily cGPS time series to get the appropriate RW (mar_neu) values for each cGPS site; use the median RW for the sGPS sites.
See sGPS_recipe.txt for detailed commands
Notes: The short-course folder also includes the GLOBK command files to be used in the recipe.

## Slide 14
SARG
DALL
BURN
217U
Locations of one continuous (BURN) and 3 survey-mode sites for time series shown in next slides
Notes: Next slides show time series for these 5 well-determined sites in E OR and WA. Two of them are continuous---BURN, we’ve seen before, and GOBS. The other three have been surveyed for 2-5 days every 1-2 years. The velocities shown here are with respect to a model (NAM) that removes the rotation of the Oregon block and elastic deformation along the boundary (diagonal from upper left corner to 45N 117W. The residuals of these sites and others near them are less than 1 mm/yr.
Next slide: BURN

## Slide 15
Time series of monthly position estimates for continuous site BURN
Wrms ~ 1 mm N,E
~ 3 mm U
Rate uncertainties
< 0.2 mm/yr N,E
0.7 mm/yr U
do not include random walk added for velocity estimates
Notes: A 5-yr time series with sampling only in the summer, at the time of field surveys, from a different processing. Here the reference frame is different, so the rate differences from the previous slide cannot be considered simple data noise. Note that the formal uncertainties from yearly samples are nearly as small as from the continuous time series: 0.2 mm/yr in horizontal, 0.7 mm/yr in vertical. You could argue here that the persistent slope in the last half of the vertical time series suggests systematic errors in vertical as large as 2 mm/yr—still close to a useful result for constraining some interesting phenomena.

## Slide 16
SARG
DALL
BURN
217U
Next slide shows time series for survey-mode site 217U
Note consistency with nearby sites
Notes: Next slides show time series for these 5 well-determined sites in E OR and WA. Two of them are continuous---BURN, we’ve seen before, and GOBS. The other three have been surveyed for 2-5 days every 1-2 years. The velocities shown here are with respect to a model (NAM) that removes the rotation of the Oregon block and elastic deformation along the boundary (diagonal from upper left corner to 45N 117W. The residuals of these sites and others near them are less than 1 mm/yr.
Next slide: BURN

## Slide 17
Time series for survey-mode site 217U
Position estimates based on 8-24 hr occupations
Note < 1 mm rate uncertainties due to 7-yr time span
Notes: With the sparcity of data here, and with DALL, you can argue that the formal errors are close to the real ones. In fact, as I’ll show later from the consistency of velocities, a normalized rms of 0.6-0.8, given four or more epochs, produces a realistic one-sigma uncertainty—here 0..5 mm/yr. This is a reliable result largely because of the long time span 7-years.

## Slide 18
SARG
DALL
BURN
217U
Next slide shows time series for survey-mode site SARG
Notes: Next slides show time series for these 5 well-determined sites in E OR and WA. Two of them are continuous---BURN, we’ve seen before, and GOBS. The other three have been surveyed for 2-5 days every 1-2 years. The velocities shown here are with respect to a model (NAM) that removes the rotation of the Oregon block and elastic deformation along the boundary (diagonal from upper left corner to 45N 117W. The residuals of these sites and others near them are less than 1 mm/yr.
Next slide: BURN

## Slide 19
Horizontal time series for survey-mode site SARG
Position estimates based on 8-24 hr occupations
Note 1 mm wrms and
< 1 mm rate sigmas
Notes: Another example of a field site which appears to be well behaved over 5 years, leading to realistic uncertainties in horizontal velocities less than 1 mm/yr.

## Slide 20
SARG
DALL
BURN
217U
Next slide shows time series for survey-mode site DALL
Note consistency with nearby sites except continuous site GWEN
Notes: Next slides show time series for these 5 well-determined sites in E OR and WA. Two of them are continuous---BURN, we’ve seen before, and GOBS. The other three have been surveyed for 2-5 days every 1-2 years. The velocities shown here are with respect to a model (NAM) that removes the rotation of the Oregon block and elastic deformation along the boundary (diagonal from upper left corner to 45N 117W. The residuals of these sites and others near them are less than 1 mm/yr.
Next slide: BURN

## Slide 21
Position estimates based on 8-24 hr occupations
Rate sigmas < 1 mm/yr and consistent with surrounding sites even with velocities determined essentially by two occupations 3 yrs apart
Horizontal time series for survey-mode site DALL
Notes: From the time series, some reason to doubt that the uncertainties are realistic. (But velocity is consistent with neighbors.)